## Slide 2
A
B
C
D

## Slide 3
This shows the prototype with the wrist, notches milled into a nitinol tube, connected to a stainless steel shaft that is clamped onto the handle that consists of a finger piece that controls the cable displacement of the cable attached to the nitinol wrist. Moving the finger piece back causes cable displacement and thus wrist actuation.
finger piece
wrist
stainless steel shaft
collet clamp
handle
cable

## Slide 5
Promontory
Sinus Tympani
This is a model of the left temporal bone. The promontory is a landmark bone inside the middle ear, behind the ossicles. The sinus tympani is shown and is very difficult to reach into with standard, rigid tools to dissect and remove cholesteatoma. Often, the cholesteatoma is visualized in the sinus tympani with the endoscope but the tools cannot reach inside to extract it. This image shows, with an endoscope view, that the controllable, flexible instrument can reach into the sinus tympani.

## Slide 6
This is a model of the left temporal bone. The model has been cropped so that the antrum is visible in this bird’s eye view. Cholesteatoma had eroded the ear canal in this patient like an atticoantrostomy, a hole in the ear canal where the instrument is coming through. Thus, the instrument is introduced through that opening and the tip can reach and dissect the boundary of the antrum.

## Slide 7
Antrum
Ossicles
Sinus tympani
Ossicles
3D virtual model of temporal bone anatomy used to identify structures for the new tool to reach.

## Slide 8
Wrist
Shaft
Thumb Piece
Handle
Collet Assembly
Laser Fibre
Laser Fibre
Cable

## Slide 9
Wrist
Shaft
Collet Assembly
Thumb Piece
Handle
Luer Lock
Connector
Tubing
Collet Clamp
Collet
Tubing
Adaptor Tube
Cable

## Slide 10
Thomassin Dissector Tip
Proposed Tool Dissector Tip

## Slide 11
Spiggle & Theis Panetti Endoscopic/Microscopic Instrument Set for Middle Ear Surgery
Suction Knife, Oval
Suction dissectors
Grace Medical IWGEES Instruments
Thomassin Dissector
Forceps
Karl Storz Middle Instrument Set for Endoscopic Middle Ear Surgery
Derlacki Mobilizer
Rosen Needle
These are different sets of endoscopic ear surgery instruments. The Spiggle & Theis Panetti set incorporates suction along the shaft of its instruments in order to allow for two functionalities, suction and dissecting or suction and cutting, in one tool, eliminating the need to switch between a suction instrument and dissection instrument or knife. Grace Medical and Karl Storz have similar suction capabilities but the PI has preferred the curves of the Panetti set. Many instruments have a curve at the tip in order to reach the middle of the endoscopic field of view, which does not occur for straight-shaft instruments as the ear canal constricts the movement of instruments. The Thomassin dissector, Derlacki Mobilizer and Rosen Needle are frequently used by the PI to position grafts and dissect tissue as their respective tip shapes are preferred to manipulate tissue effectively. The Rosen Needle tip curvature allows for dissecting tissue that is attached to the ear drum as the tip shape compliments the curvature of the ear drum, and when the needle is rotated axially, the tip follows a trajectory which is useful to manipulate tissue without having to translate the tool which would cause the instrument shaft to collide with the endoscope or ear canal wall. Other companies such as Bausch and Lomb and Sklar also have endoscopic ear surgery instruments. http://www.spiggle-theis.com/en/products/otology/panetti-instrument-set
http://www.gracemedical.com/products/endoscopic-instruments/endoscopic-instruments/
https://www.karlstorz.com/cps/rde/xbcr/karlstorz_assets/ASSETS/3448115.pdf

## Slide 12
Steerable Instruments
This shows current instruments that are steerable. The vitreoretinal laser probes are used to deliver laser energy to the retina for therapy. The slider at the handle allows the surgeon to push out the preshaped 90o nitinol tip. The surgical instrument is a manually engageable handle + steerable tip with pin joints.
http://salientmed.com/solution/endoprobe-handpieces/
https://patents.google.com/patent/US7766904B2/en?q=endoprobe&q=adjustable&q=flexible https://www.google.com/patents/US5454827
Vitreoretinal laser probes
Surgical Instrument